CNSI Microfluidics Lab - Harrick PDC-32G Plasma Cleaner Operations

This procedure is adapted from the University of Arizona’s SOP.

This machine is to be used by authorized personnel only.

Load Sample:

- Ensure that your sample fits inside the plasma chamber.

- If present, remove the glass sample holder plate from chamber carefully.

- Place your sample onto the mid-section of the glass holder plate (if available).

- Insert the glass holder plate and your sample back into plasma chamber carefully.

- Ensure vent valve on the chamber lid is closed.

- Ensure rubber O-ring on the lid is clean and free of any debris.

- Place lid over the chamber opening and hold lid in place with one hand.

- Turn on the PUMP ON/OFF switch to turn on the vacuum pump.

- Chamber vacuum should now hold lid in place.

Run Plasma:

- Ensure RF LEVEL power switch on the front panel is on OFF.

- Turn on the POWER switch on the front panel.

- Wait for one minute before proceeding.

- Open the chamber lid vent valve about 1/8 of a turn (counter-clockwise) to establish an air flow into the chamber. Note: some users find more consistent results with a Nitrogen plasma instead of flowing air in the chamber. If you would like to do this:

  - Open the Nitrogen valve on the right side of the fume hood

  - Close the valve at the bottom of the left-hand flow meter on the plasma cleaner controller.

  - Open the chamber vent valve about 1/8 of a turn

  - Open the flow meter valve slightly until the chamber pressure is 450mT

- Switch the RF LEVEL power switch to the desired setting.

- Look for a purple glow which would indicate the presence of a plasma.

- Run plasma for the desired time using an external timer.

Unload Sample:

- Switch the RF LEVEL power to OFF.

- Turn off Nitrogen gas valve on the hood

- Turn off the POWER switch.

- Hold the lid until chamber is fully vented and the lid comes off.

- Open the chamber lid vent valve slowly until you hear hissing sounds from the valve to vent chamber completely.

- Turn off the PUMP ON/OFF switch to turn off the vacuum pump. Start chamber vent before turning pump off.

- Remove the glass sample holder from chamber and remove sample.

- Clean the glass sample holder if necessary and re-insert back into chamber.