## Slide 1: Training Chromeleon 7.2.10 – QqQ – Small Molecules
LSMS

## Slide 2: Summary
Module 1: Starting Chromeleon 7.2 workstation
Module 2: Chromeleon Console
Module 2.1: Chromeleon Console: Administrative Tasks
Module 2.2: Chromeleon Console: Data Vault
Module 2.3: Chromeleon Console: Data Category Bar
Module 2.4: Chromeleon Console: The Sequence
Module 2.5: Chromeleon Console: The Injection List and Associated Items
Module 2.6: Chromeleon Console: Instrument Category Bar
Module 2.7: Chromeleon Console: eWorkflows Category Bar
Module 3: Data Acquisition
Module 3.1: Creating a Sequence
Module 3.2: Creating an Instrument Method
Module 3.3: Creating a Shutdown Instrument Method
Module 3.4: Creating a Processing Method
Module 3.5: Running the Sequence
Module 4: Chromeleon Studio
Module 4.1: Chromeleon Studio - Chromatogram and Mass Spectra
Module 5: Building Processing Method
Module 5.1: Data Processing - Processing Method / MS Component Table
Module 5.2: Data Processing - MS Component Window
Module 5.3: Data Processing – Calibration Levels
Module 5.4: Data Processing - Properties Component
Module 5.5: Data Processing - MS Components
Module 5.6: Data Processing – Calibration Plot
Module 5.7: Data Processing – Interactive Results
Module 5.8: Data Processing – Interactive Charts
Module 5.9: Data Processing - Using SmartLink in MS Components window
Module 6: Report Templates
Module 7: Creating an eWorkflow
Module 8: Exporting and Importing Data

## Slide 3: Chromeleon Architecture
The Chromeleon 7.2 workstation is a client/server program, which allows operation of the data system locally or in a network.
Client: The Chromeleon program is referred to as the client.
  The user can perform tasks such as creating Instrument and Processing Methods, viewing data, searching for samples, etc.
Server (Instrument Controller): The Instrument Controller PC is connected directly to the modules.
  The Instrument Controller executes the active control or communication to the modules that are powered on and configured to the Instrument.

## Slide 4: Module 1: Starting Chromeleon 7.2 workstation
Module 1: Starting Chromeleon 7.2 workstation

## Slide 5: Starting Chromeleon 7.2 workstation
Since the Client and the Server (Instrument Controller) are independent programs they must be started separately.

## Slide 6: Starting Chromeleon 7.2 workstation
The Instrument Controller and Chromeleon Client applications can be started in any order, but on a Chromeleon laboratory workstation, it is recommended to start the Instrument Controller before launching the client.

## Slide 7: Checking the Status of the Instrument Controller
First, check the status of the Instrument Controller by hovering the cursor over the Instrument Controller icon (the Charlie icon, indicated by the red box) in the notification area of the Windows taskbar.
NOTE: If the Services Manager icon does not appear in the Windows System Tray, then it may be started by clicking the Windows Start icon and selecting All Programs > Thermo Chromeleon 7 > Services Manager.

## Slide 8: Checking the Status of the Instrument Controller
A small box appears and indicates the status of the Instrument Controller.
  If the box reads “The Instrument Controller is stopped” or the Charlie icon has an “X” on it (as shown above) then it needs to be started.

## Slide 9: Starting the Instrument Controller
To start the Instrument Controller:
  Right-click on the Chromeleon icon (“Charlie”) and choose the “Show Chromeleon Service Manager” option.

## Slide 10: Starting the Instrument Controller
  The Chromeleon Services Manager window opens.
  Click “Start Instrument Controller”.
  NOTE: You can check the “Start service on system start” so that the Instrument Controller automatically starts when the computer starts up.

## Slide 11: Starting the Instrument Controller
When “Stop Instrument Controller” appears (as shown above), close the Chromeleon Services Manager pop-up window.
  The “X” will be removed from the “Charlie” icon.
NOTE: This step is only required when the Instrument Controller PC is connected directly to the modules, and as a daily operator, this is your only interaction with the Chromeleon Services Manager. If running Chromeleon 7.2 on a standalone desktop, you do not have to start the Instrument Controller.

## Slide 12: Chromeleon Instrument Configuration
Before the Chromeleon 7.2 workstation can run the instrument, it has to be configured.
  This is typically done during installation.
NOTE: You can configure the instrument by right-clicking on the Chromeleon icon (“Charlie”) on the task bar and choosing “Configure instruments” option.

## Slide 13: Creating an Instrument
In the Instrument Configuration window that appears, click on the computer name (FRLES-9TFR4M2), right-click and choose the “Add Instrument” option.
  Give the Instrument a name.

## Slide 14: Creating a Module
Next, right-click on the Instrument, and choose the “Add Module” option.

## Slide 15: Creating a Module: Pump
In the pop-up window that appears, choose your Manufacture and Module you want to add. In this example, we chose “HPLC Vanquish” and “Vanquish Binary Pump”.
Click “OK”.

## Slide 16: Creating a Module: Pump
If you are not connected to an instrument, ensure you select the “Enable Simulation Mode”.
However, if you are connected to an instrument, click “Browse” to get the IP address for the component.

## Slide 17: Creating a Module: Pump
If you change the Pressure Unit from “bar” to “psi”, ensure you update the methods to read the new Pressure Units.

## Slide 18: Creating a Module: Pump
NOTE: You can name your solvents under the Solvents tab.

## Slide 19: Creating a Module: Column, Autosampler & DAD
Also, add Vanquish:
  Column Compartment
  Autosampler
  Diode Array Detector (if needed)

## Slide 20: Editing a Module
After adding all the Modules, you can edit a specific Module by right-clicking on it (e.g. the Column Compartment) and choosing Properties...

## Slide 21: Editing a Module
In this example, we are editing the Column Compartment.
  Hence the Column Compartment Configuration window appears.

## Slide 22: Editing a Module
Here, the column compartment is on the right, thus we are only using the Right Pre-heater.
To not use the Left Pre-heater, uncheck it.
NOTE: Remove the voltage connection (there’s still a latent heat of ~30 ºC when plugged in).

## Slide 23: Creating a Module: MS
Add another module, choose Manufactures: “Mass Spectrometry” and Modules: “Mass Spectrometer” for QE or TSQ.

## Slide 24: Configuring MS communication
When the MS is connected to the Instrument Controller, the Instrument identification and IP address will be automatically populated (red box).

## Slide 25: Saving the Instrument configuration
Click File > Save Installation.
Close the Instrument configuration window.
NOTE: The number of instruments you can control is dependent on your license.

## Slide 26: Saving the Instrument configuration
NOTE: By default, the .cmic (Chromeleon Instrument Configuration) file (blue box) is stored in the location highlighted above (red box).

## Slide 27: Backing up the Instrument Configuration
You can backup the instrument configuration, by going to File > Export, which allows you to access it even after a system restore.
NOTE: The exported .cmic (Chromeleon Instrument Configuration) file can be stored in any location, e.g. the Documents folder shown above.

## Slide 28: Starting the Chromeleon client
To start the Chromeleon client, double-click on the Chromeleon Windows Desktop icon.
  The Console is the first window to appear upon starting the Chromeleon 7.2 client.

## Slide 29: Module 2: Chromeleon Console
Module 2: Chromeleon Console

## Slide 30: The Chromeleon Console
Within the Console window the user can control Instruments, navigate through all data and use and manage eWorkflows.

## Slide 31: Layout of the Chromeleon Console
The Chromeleon Console is divided into three main sections: Category Bars, Navigation Pane, and Work Area.
Work Area
Navigation Pane
Category Bars

## Slide 32: Chromeleon Console: The Category Bars
Category Bars are used to switch between Instrument control, Data review, and eWorkflows.
  You click on the corresponding Category name to change the view.

## Slide 33: Chromeleon Console: The Navigation Pane
The Navigation Pane is an area that allows navigation through objects related to the selected category.

## Slide 34: Chromeleon Console: The Work Area
The Work Area is an area that displays the content of the object selected in the Navigation Pane, and allows users to work with that object.

## Slide 35: Module 2.1: Chromeleon Console: Administrative Tasks
Module 2.1: Chromeleon Console: Administrative Tasks

## Slide 36: Accessing the Administration Console
To access the Administration Console:
  Go to Tools → Administration Console…

## Slide 37: Client Licenses
In the pop-up window, choose the License Manager (localhost) option in the left navigation pane.
Next, click on the “Client Licenses” hyperlink in the right navigation pane.

## Slide 38: Client Licenses
The Client Licenses window should appear as shown above.

## Slide 39: Client Licenses
Choose your “Consumer Type” & “License Consumer”.

## Slide 40: Client Licenses: Compliance Tools license
If you have the additional Compliance Tools license, click the “Compliance Tools” option.
  This enables user management, client locking, versioning of data objects, and data audit trails.
NOTE: While Chromeleon allows you to track changes to data with the implementation of the Compliance Tools license by saving different “versions”, these “versions” are not copies of the data.

## Slide 41: Client Licenses: Report Design Pro license
NOTE: Spreadsheet cells have the capability of performing Excel‐type functions such as IF/THEN, AVERAGE, STDEV, and more.
However, this requires the Report Designer Pro license. Newer versions of Chromeleon 7.2 (SR6 and later) all come with this license.

## Slide 42: Instrument Controller Licenses: LC-MS Data Acquisition
Clicking the “Instrument Controller Licenses” option in the left pane allows you to see the number of Instrument Controller Licenses being used.
NOTE: Checking the “LC-MS Data Acquisition” option is required for operation of an MSQ, TSQ Endura / Quantiva or Exactive MS driver.

## Slide 43: Instrument Controller Licenses: 3D Data Acquisition
Select “3D Data Acquisition” to enable recording of 3D data e.g. UV data for an instrument controller.

## Slide 44: Instrument Controller License Status
If an instrument controller is running, its current license status is shown in the Status column. If the instrument controller is offline, the last known status is shown.
  OK (green color): The license configuration matches the requirements of the instrument controller.
  New (yellow color): No licenses have been configured for the instrument controller yet.
  Update (red color): There is a mismatch between the license configuration and the requirements of the instrument controller.
You may have to update your Chromeleon license to fully support the instrument controller's requirements.
If the status is unknown, the status field remains empty.

## Slide 45: Instrument Class 1, 2 & 3
NOTE: There are three Instrument Classes (each one can control a different number of instruments).
All Thermo Scientific and Dionex IC, LC, and GC instruments, including ESA and LC Packings modules, can be controlled with an Instrument Class 1 license.
All third-party GC systems can be controlled with an Instrument Class 2 license.
All third-party LC systems can be controlled with an Instrument Class 3 license.
  Class 3 licenses can control Class 2 & Class 1 instruments, while Class 2 licenses can control Class 1 instruments.

## Slide 46: Instrument Class 1, 2 & 3 License Status
The color of a cell indicates the license status:
  If more licenses are requested than can be granted, the related cell has a red background.
  If fewer licenses are requested than granted, the related cell has a yellow background.
  If the number of license features granted matches the number of licenses requested, the cell has a white background.

## Slide 47: License Overview
As stated previously, the number of instruments you can control is dependent on your license.
To see the total number of Client or Instrument Controller licenses, select the “License Overview” option in the left pane.

## Slide 48: Creating Users
You can also create new Users. To do this:
  Go to Tools → Administration Console…

## Slide 49: Creating Users
In the pop-up window, select Global Policies in the left navigation pane, and User Mode (under the List of policies list).
Check “Enable User Mode”.
In the “Confirm Enable User Mode” window, input a User Name and Password.

## Slide 50: Scheduling Tasks
We can schedule Administrative tasks by going to Tools → Administration Console…
In the pop-up window, select Scheduler in the left navigation pane.
  We can schedule a task such as backing up data either to another computer/server that has Chromeleon installed on it or locally.

## Slide 51: Scheduling Data Backup
To schedule data back up on another computer/server, click on the “Scheduler (Chromeleon Domain)” option.
  In the Task List tab that appears, click “Create new task” (red box).

## Slide 52: Scheduling Data Backup
In the Create New Task wizard, in the General tab, give the task a Name.
Next, go to the Schedule tab, and input how frequently you want to perform the task.

## Slide 53: Scheduling Data Backup
In the Action tab, click “New”.
In the “Create New Task Action” pop-up window that appears, chooses your action: “Copy”, “Move” or “External Program”.
Click the arrow next to Data vault to select the destination data vault.

## Slide 54: Scheduling Data Backup
Next, click the icon in the Folder formula: tab.
In the “Insert Formula” pop-up window that appears, select “Sequence” in the Categories section and “Data Vault” in the Variables section.
  This creates a folder that is named after the Data Vault assigned to the sequence.
Click “OK” (red box).
Next, click “OK” (blue box) in the “Create New Task Action” pop-up window.

## Slide 55: Scheduling Data Backup
Click “Finish” (red box) in the “Create New Task Wizard” pop-up window.

## Slide 56: Module 2.2: Chromeleon Console: Data Vault
Module 2.2: Chromeleon Console: Data Vault

## Slide 57: The Data Vault
Data and all associated items are accessed from the Chromeleon Console’s Data Category.
Objects are organized and stored in Data Vaults, a container for all Chromeleon data and objects.
A default installation of Chromeleon 7 contains a local Chromeleon 7 Data Vault named “ChromeleonLocal”, which is stored on the C: drive.

## Slide 58: Storing data elsewhere
You can store the data elsewhere by creating a new Data Vault. To do this:
  In the Menu bar, click Tools, and select Administration Console…

## Slide 59: Creating a new Data Vault
In the pop-up window that appears, select Local Machine, and click “Manage Data Vaults”.

## Slide 60: Creating a new Data Vault
In the Data Vault Manager – Local Machine window that appears, click “Create”.
Select “Create standard data vault” in the Create Data Vault window.
Click “Next”.
If data storage is not enough, you can use Enterprise solution (this is analogous to a workgroup).
It stores data temporarily in Xvault, and when connected to the network, it syncs with the remote server.

## Slide 61: Creating a new Data Vault
Give it a name.
Check “Enable versioning and data audit trails”
  You can restore to previous versions.
Choose Data vault path.
Click “Finish”.

## Slide 62: Chromeleon Tips & Tricks: Communication Issues
If the Data Vault is not connected (red X next to the Data Vault in the Navigation pane):
  Go to the Windows Start Menu, and type “Services”.
  Click Services under Service Manager.

## Slide 63: Chromeleon Tips & Tricks: Communication Issues
In the pop-up window, select the “SQL Server”, and Restart it.

## Slide 64: Chromeleon Tips & Tricks: Communication Issues
Alternatively, you can mount and dismount the Data Vault. To do this:
  In the Chromeleon Console, go to Tools > Administration Console…
  Select Local Machine
  Select Manage Data Vaults
  Select the Data Vault that has the issue, and either click “Mount” or “Dismount” based on which you are doing.
  Select Refresh.

## Slide 65: Creating a folder
You can also create new Folders. To do this:
  Select the folder in which the new folder should be created.
  In the Menu Bar, click Create and choose Folder.

## Slide 66: Creating a folder
Give the new folder a name.

## Slide 67: Moving Objects Between Folders: “Drag and Drop”
Objects such as Instrument Methods, Processing Methods, View Settings, and Report Templates can be copied and moved between folders using the “Drag and Drop” feature.
To drag and drop:
  Click and hold the left mouse button on the object of interest.
  Move the object to the desired location.

## Slide 68: Moving Objects Between Folders: “Drag and Drop”
If the dialog box shown above appears, choose Copy to make a copy of the original object in the target location.
Or choose Move to move the original file to the new location.
NOTE: Copies of objects are independent of the objects from which they were copied. For example, shortening the analysis time of an Instrument Method within a folder has no effect on the copied file in another folder.

## Slide 69: Moving Objects Between Folders: “Drag and Drop”
The default behavior for Drag & Drop may be set by selecting from the Menu bar Tools → Preferences.
The Preferences are as follows:
  Ask: User is prompted to confirm each Drag & Drop action within a Data Vault.
  Copy: The selected element is copied to the new position without being deleted from the original position.
  Move: Within a Data Vault: The selected element is moved to the new position and deleted from the original position. To a different Data Vault: The selected element is copied to the target position, but is not deleted from the original location.

## Slide 70: Moving Objects Between Folders: “Right-click” or “Edit” menu
Alternatively, you can right‐click on an object to open the Context menu and select Copy, Cut or Paste commands or select these commands from the Edit menu.

## Slide 71: Locking folders
Folders may be locked to prevent the accidental deletion or modification of objects and data.
  This does not prevent a user from unlocking the folder and making changes or deletions.
To lock a folder or sequence, right‐click on the folder or sequence and select Read‐Only.
NOTE: If the additional Compliance Tools license is not in place you can not lock folders.

## Slide 72: Locking folders
When a top‐level folder is locked then all sub‐folders and Sequences are also locked.
Locked Top-level folder
Locked sub-folders

## Slide 73: Module 2.3: Chromeleon Console: Data Category Bar
Module 2.3: Chromeleon Console: Data Category Bar

## Slide 74: Accessing Sequences or Objects in the Data Category Bar
The Data Category provides access to data and objects used to acquire, process and report data.
When the Data Category is selected on the Navigation Pane the user can select Sequences or Objects stored within the Data Vaults.
Sequences or Objects
Data Category selected

## Slide 75: Accessing Data & Objects in the Data Category Bar
By clicking on a sequence you can see the data, along with all the associated items for that sequence in the Work Area.
NOTE: Though the structure of the Chromeleon Console is similar to that of Windows Explorer, Chromeleon data cannot be accessed through Windows Explorer.
Objects:
Instrument & Processing Methods
Report Templates
View Settings
Data

## Slide 76: Module 2.4: Chromeleon Console: The Sequence
Module 2.4: Chromeleon Console: The Sequence

## Slide 77: The Sequence gives us access to the Injection List
A list of samples is required for the uninterrupted analysis of multiple samples on the LC-MS instrument.
This list is called the Injection List and is contained in the Sequence.
  Sequences appear in the Navigation Pane as two autosampler vials.
Injection List
Sequence

## Slide 78: Sequence contains all items required to acquire & process data
When a Sequence is selected in the Navigation Pane, the Injection List and Associated Items are displayed in the Work Area.
  Hence, the Sequence contains all items required to run samples and to process, view and report the acquired data.
Associated Items
Injection List
Sequence

## Slide 79: The Injection List can be used to view raw data in real-time
The raw data is saved in the Sequence as the analysis progresses and may be accessed (in real‐time) by viewing “Mini Plots” from the individual injections in the Injection List.
Mini Plots (red box) are thumbnail chromatograms that allow the user to quickly view the data.

## Slide 80: Mini plots can show any data analysis channel acquired
When multiple channels are acquired during an analytical Sequence, the data analysis channel displayed on the Mini Plots may be changed by clicking on the right arrow button at the top of the Mini Plot column.
NOTE: Mini Plots are scaled according to the View Setting that is currently being used in the Chromatography Studio (discussed later) and are even available for the currently running analysis.

## Slide 81: Sequence Control Bar
The Sequence Control Bar is located in the Work Area above the Injection List, and may be used to perform such tasks as viewing the Sequence state, start or stop the Sequence, or viewing the Instrument status.
Sequence state
Start or Stop the Sequence
Viewing the Instrument status

## Slide 82: Sequence State
Sequence state

## Slide 83: Instrument Status
Viewing the Instrument status
The different actions in the Instrument Status will affect only the LC, not the MS.

## Slide 84: Module 2.5: Chromeleon Console: The Injection List and Associated Items
Module 2.5: Chromeleon Console: The Injection List and Associated Items

## Slide 85: The Injection List (in more detail)
Each line represents an injection and each column contains properties of the injection, including the name of injection, type of injection (Blank, Calibration Standard, Unknown), Injection Volume, location in the autosampler, Dilution Factors (if any), etc.
In addition, each injection must be assigned an Instrument Method and Processing Method.
  Instrument Methods are created to control hardware operation.
  Processing Methods contain data processing parameters.

## Slide 86: The Injection List: Injection Type
The Injection Type indicates to the software the purpose of the Injection.
The Injection Type is assigned by left‐clicking on the box and then using the pull‐down menu to select the appropriate type.

## Slide 87: The Injection List: Injection Type
There are different Injection Types indicated by icons next to the Injection Name.

## Slide 88: Associated Items (in more detail)
The lower section of the Work Area of the selected Sequence contains various Associated Items used when running the Sequence or for the interpretation of the acquired data.
  All items the Sequence can utilize are listed in this section.
Each object type is identified by an icon, and its type is listed next to its name in the Work Area.

## Slide 89: Module 2.6: Chromeleon Console: Instrument Category Bar
Module 2.6: Chromeleon Console: Instrument Category Bar

## Slide 90: Instruments Category Bar
The Instruments Category provides immediate access to and control of all available Instruments.

## Slide 91: Filtering Instruments available in the Chromeleon Console
When the Instruments Category is selected the user can select on the Navigation Pane an Instrument to monitor and/or control.
  When the icon is clicked, you can see all the local instruments.
  When the icon is clicked, you can see all the instruments on a network
  If you are on a network that has many instruments across various sites, you can assign your instrument as a favorite instrument by clicking on the next to it.
  You can then see all the favorites by clicking the icon at the top.

## Slide 92: Filtering Instruments available in the Chromeleon Console
Click on the instrument you want to use.
The Work Area then immediately displays the ePanels for the selected Instrument.

## Slide 93: ePanels that do not depend on instrument configuration
No matter the instrument configuration, the following ePanels are always present:
  Home
  Audit
  Queue

## Slide 94: Home ePanel
The Home ePanel gives you access to all modules configured for the instrument in one location.
  You can do more advanced tasks for each Module in its specific ePanel.

## Slide 95: Audit ePanel
The Audit ePanel contains the Instrument Audit Trail, which is a daily event log that records all events related to instrument operation, such as system events, pre-run device settings, executed commands, and error messages.
  You can also see it in Real Time or as “History” over a previous period of time.

## Slide 96: Queue ePanel
The Queue ePanel provides controls for managing the sequence Queue.

## Slide 97: Module 2.7: Chromeleon Console: eWorkflows Category Bar
Module 2.7: Chromeleon Console: eWorkflows Category Bar

## Slide 98: eWorkflows Category Bar
An eWorkflow contains a set of templates and a set of rules. It is the easiest way to build an error-free Sequence.
  The rules define, for example, the Sequence layout and the list of available Instruments that are suitable for the analysis.
When you start an eWorkflow a Sequence is created with associated objects and a well‐defined structure.

## Slide 99: eWorkflows Category Bar
When the eWorkflows Category is selected the user can select on the Navigation Pane an eWorkflow to execute.
  The Work Area then displays the Instruments available to run the eWorkflow.

## Slide 100: eWorkflows Category Bar
After selecting the instrument, click Launch.
  This opens the sequence.
Click “Finish”.

## Slide 101: eWorkflows Category Bar
The Sequence opens in the Data Category Bar.
Click “Start” to run the sequence.
NOTE: This eWorkflow generates a folder name based on the instrument used and a sequence name based on the time it was generated (red box).

## Slide 102: eWorkflows Category Bar
NOTE: There are two things you need to know before creating an eWorkflow:
Order of Sequence
Instrument Method, Processing Method, Report Template and View Settings.

## Slide 103: Module 3: Data Acquisition
Module 3: Data Acquisition

## Slide 104: First Steps…
Before we acquire data, we first need to create a Sequence, followed by creating Instrument and Processing methods.
NOTE: Under the Create tab, the tasks are listed in a logical way of how you would run a workflow from top to bottom.

## Slide 105: Module 3.1: Creating a Sequence
Module 3.1: Creating a Sequence

## Slide 106: Creating a new sequence
Select the folder where you want the data to be stored.
Next click Create > Sequence…

## Slide 107: Creating a new sequence
Choose the Instrument, and then go through the Wizard.

## Slide 108: Setting up injection list
Input an Injection Name, the number of vials, the number of replicates, vial position and injection volume.
  A Sequence Preview is shown at the bottom of the window.

## Slide 109: Setting up injection list
NOTE: You can click the arrow next to the Injection Name and choose one of the options.

## Slide 110: Setting up injection list
E.g. The end of the Injection name has changed with the replicate number for each sample (red box).

## Slide 111: Adding Methods and Defaults
If you have an Instrument Method, Processing Method, Report Template or View Setting, you can add them here.
  If you save the Report Templates or View Settings as a link, any change made to them in the current sequence will affect the original.

## Slide 112: Adding Methods and Defaults
If you don’t have an Instrument Method, Processing Method, Report Template or View Setting, leave blank.
Click Next.

## Slide 113: Features
You can use this option to determine the settings used for updating views, reports, or charts for MS data in the Chromeleon Studio.
NOTE: The Modify when Visible Views are updated privilege (found in the Privileged Actions in the Global Policies section of the Administration Console) is required to modify this option.

## Slide 114: Features
Automatically (default): When selected, views, reports and charts for MS data are updated automatically each time processing parameters are changed.
On-demand: When selected, views, reports and charts for MS data are updated only when you click Ribbon Bar.
Click Next.

## Slide 115: Saving the new sequence
You can then add a Comment.
Click Finish.
In the pop-up window, input a name and click Save.

## Slide 116: Viewing the new sequence
The new Sequence should now appear in the Navigation Pane along with the Injection List in the Work Area.

## Slide 117: Editing the new sequence
You can now input the correct sample names by editing the Injection List in the Work Area.

## Slide 118: Deleting an existing Sequence
You can delete a sequence by selecting it, right-clicking and choosing the “Delete” option.

## Slide 119: Restoring a deleted Sequence
The deleted Sequence should now be in the Trash.
You can either Restore or completely remove the deleted Sequence by selecting it, right-clicking and choosing the appropriate option.
NOTE: Choosing “Restore” option, restores the Sequence to the original folder, whereas the “Restore To…” allows you to restore it to another folder.

## Slide 120: Module 3.2: Creating an Instrument Method
Module 3.2: Creating an Instrument Method

## Slide 121: Creating a new Instrument Method
If you haven’t created an Instrument Method, go to the Navigation Pane and select the Sequence you created earlier.
Next on the ribbon at the top, click “Create” > “Instrument Method”.

## Slide 122: Instrument Method – Select Instrument
In the Instrument Method Wizard, select the Instrument.
Click Next.

## Slide 123: Instrument Method - Overview
In the Instrument Method Wizard, enter the Run Time and select the Diagnostic Channels you want to monitor.
Click Next.

## Slide 124: Instrument Method - Pump
In the next section of the Instrument Method Wizard, you can input the Solvent Names and specify the Solvent that will be used (%A1 or %A2 or %A3) and (%B1 or %B2 or %B3).
Click Next.

## Slide 125: Instrument Method – Flow Gradient
In the Flow Gradient section of the Instrument Method Wizard, first click “Remove Equilibration Stage” (red box).
Click “Yes” in the pop-up window.

## Slide 126: Instrument Method – Flow Gradient
Edit your gradient.
Click Next.

## Slide 127: Instrument Method – Sampler Module
We typically use a Wash mode of “Both”, which washes the Needle before and after sample aspiration.
Click Next.

## Slide 128: Instrument Method – Sampler Module Temperature
We can set a temperature for the sampler module by clicking “Use temperature control” and entering a temperature.
Click Next.

## Slide 129: Instrument Method – Column Compartment
Enable temperature control for the Column chamber and Pre-Heater; then enter the temperature you will use.
NOTE: Ensure you select “Advanced” (red box) to see all the parameters.
Click Next.

## Slide 130: “Forced Air” vs. “Still Air”
Forced Air
  Backpressure (≥ 400 bar)
  Solvent pre-heated
  Reduced axial temperature gradient
  Self-heating by friction
  Radial temperature gradient
  Low axial temperature gradient
  Peak broadening
Still Air
  Backpressure (≥ 400 bar)
  Solvent pre-heated
  Reduced axial temperature gradient
  Self-heating by friction
  Almost no radial temp. gradient
  Axial temperature gradient
  Sharp peak
70 °C
90°C
70 °C
70 °C
80 °C
73 °C
76 °C
70 °C
Flow concentration
Same temperature at column surface
Notes: There is always frictional heating that occurs in your column, so there is always a temperature gradient along the column i.e. the temp at the outlet is higher than the inlet.
There is also a radial gradient (temp on the inside is hotter than the outside).
At high pressures:
In forced air mode, frictional heating results in an increase in temp. The air at the walls of the column then returns to the column heater. The peltier then compensates for the hotter air by decreasing the temp, which results in cold spots in the column oven (and cooler temperatures at the surface of the column)
- Peak broadening results due to molecules having a faster linear velocity at the center of the column.

## Slide 131: Above 400 bar use “Still Air”
For Column Chamber, there is no difference between Forced Air or Still Air mode below 400 bar.
  You can do Forced Air during equilibration, as it heats faster; but in the instrument method run samples in Still Air mode.
Above 400 bar use “Still Air”.

## Slide 132: Instrument Method – Column Compartment
NOTE: If you have a chip on your column, the data read from that chip will be shown under this tab.
Click Next.

## Slide 133: Instrument Method – DAD
If you have a configured DAD, specify the wavelengths.
Click Next.

## Slide 134: Instrument Method – MS Device
In Global Parameters, define the source conditions and external hardwares used.

## Slide 135: Instrument Method – MS Device
In Scan Parameters, define the aquisition parameters.

## Slide 136: Instrument Method – Comment
You can enter a Comment in this section of the Instrument Method Wizard as the used column.
Click Finish.

## Slide 137: Instrument Method – Overview
The new Instrument Method should automatically open in the Chromeleon Studio (discussed in detail later).

## Slide 138: Instrument Method – Overview
Use the Check Method option to be sure that the method is successful (red box).

## Slide 139: Saving the Instrument Method
Next, go to the icon, and select Save.
In the pop-up window that appears, input a Name and click Save.

## Slide 140: Saving the Instrument Method
The Instrument Methods you just created should now appear in the Associated Items section of the Work Area (red box).

## Slide 141: Input the Instrument Method
Next, select the Instrument Method category for Injection 1, and click on the drop down arrow.
Choose the appropriate Instrument Method.

## Slide 142: Input the Instrument Method
Assign the appropriate Instrument Method to all the injections.
You can use the F9 button on your keyboard to fill down and populate all injections that have the same Instrument Method.

## Slide 143: Module 3.3: Creating a Shutdown Instrument Method
Module 3.3: Creating a Shutdown Instrument Method

## Slide 144: Instrument Method – Smart Startup Shutdown
NOTE: The Smart Shutdown/Standby tab in an Instrument Method only applies to the LC. Hence, we need to create a Shutdown method that will put the MS into Standby.

## Slide 145: Creating a Shutdown Instrument Method
Copy/Paste and Rename the instrument method you created
Rename the instrument method : InstrumentMethod_StandBy

## Slide 146: Shutdown Method – Script Editor
Open the InstrumentMethod_StandBy and go to Scrip Editor
To put the MS in Standby, we need to use the Script Editor.

## Slide 147: Script Editor: Insert Command
Go to the line marked “End”, right-click and choose Insert Command.

## Slide 148: Script Editor: Choose Delay Command
Start typing “Delay” in the Command column, and choose the “Delay” option.

## Slide 149: Script Editor: Input Delay Value
Input a Delay value of 10 s.

## Slide 150: Script Editor: Insert Command
Go to the line marked “End”, right-click and choose Insert Command.

## Slide 151: Script Editor: Choose MSDevice → OperatingMode Command
Click the drop-down arrow (indicated by the red arrow) in the Command column, and select MSDevice → OperatingMode.

## Slide 152: Script Editor: Input Standby Value
icon should appear. Go to the Value column, click the drop-down arrow, select “Standby” and save the method.

## Slide 153: Input the Shutdown Instrument Method
You can add a new injection by going to the last line in the Injection list (red box) and either clicking on “Click here to add a new injection” or pressing the down arrow.

## Slide 154: Input the Shutdown Instrument Method
Rename the new injection and define the standby method as the new instrument method.
Next, ensure the Sample position is correct.

## Slide 155: Adding a Blank injection
Next, click the Save icon (red box) to save the sequence.

## Slide 156: Module 3.4: Creating a Processing Method
Module 3.4: Creating a Processing Method

## Slide 157: Creating a new Processing Method
If you haven’t created a Processing Method, go to the Navigation Pane and select the Sequence you created earlier.
Next on the ribbon at the top, click “Create” > “Processing Method”.

## Slide 158: Creating a new Processing Method
In the pop-up window that appears, choose “MS Quantitative”.
Click “Next>>”.

## Slide 159: Creating a new Processing Method
Input a Name.
For the “Assign the new Processing Method to” option, choose “all injection of the sequence”.
Click “Finish”.

## Slide 160: The Processing Method in the Chromeleon Studio
The new Processing Method should automatically open to the Data Processing tab in the Chromeleon Studio.

## Slide 161: The Processing Method in the Chromeleon Console
If you return to the Chromeleon Console, the new Processing Method should be assigned to all the injections (red box) and also appear in the Associated Items section of the Work Area (blue box).

## Slide 162: Module 3.5: Running the Sequence
Module 3.5: Running the Sequence

## Slide 163: Starting the Sequence
In the Status bar, click Start.
If any changes were made to the sequence and not saved, the pop-up window shown above (red box) appears.
Click “Save”.

## Slide 164: Running the Sequence
If the Ready Check for the Sequence passes, then the Sequence starts.
NOTE: The Status has changed from “Queued” to “Running” in the Status bar.

## Slide 165: Adding to the queue
Alternatively, you can click the drop-down arrow next to the Start icon in the Status bar.
Choose the “Add to Queue” option.

## Slide 166: Viewing to the queue
Again, click the drop-down arrow next to the Start icon in the Status bar.
Choose the “View Queue” option.
  This allows you to see if there are any pending Sequences.
NOTE: The Status has changed from “New” to “Queued” in the Status bar.

## Slide 167: Viewing to the queue
The Queue ePanel should appear.
  An ePanel set is a group of ePanels (modules) where you monitor and control a module's operation

## Slide 168: The Queue ePanel: Ready Check
Click Ready Check (red box) to check the Sequence.
The results are shown at the bottom (blue box).

## Slide 169: The Queue ePanel: Start Sequence
If the Ready Check is successful, click “Start”.

## Slide 170: The Queue ePanel: Sequence Status
The Status should now show that it is Running.

## Slide 171: Completed Sequence
NOTE: When the run is complete, all the runs in the Sequence show “Finished” in the Status column in Chromeleon Console.
Additionally, the status changes from “Running” to “Finished” in the Status bar.

## Slide 172: Re-running all injections in the Sequence Tips & Tricks
You can change the Status from “Finished” to “Idle” (red box) so that you can re-run that line again in the same sequence and press F9 to re-run all the sequence (blue box).
  This deletes the data acquired for the previous run permanently.
NOTE: The Status bar at the top changes from “Run Finished” to “Incomplete” and from “Incomplete” to “New”.

## Slide 173: Re-running all injections in the Sequence Tips & Tricks
To save a copy of the edited Sequence, choose the File > Save As option.
  You can save the data by choosing the “Save Raw Data” option in the pop-up window that appears.

## Slide 174: Re-running all injections in the Sequence Tips & Tricks
The new Sequence should now appear.
NOTE: The status of all Injections are “Idle” and there are no Inject Times for any of the samples in the Sequence.

## Slide 175: Module 4: Chromeleon Studio
Module 4: Chromeleon Studio

## Slide 176: Accessing Chromeleon Studio
To access the Chromeleon Studio, click on the icon or double-click on a run.

## Slide 177: Module 4.1: Chromeleon Studio - Chromatogram and Mass Spectra
Module 4.1: Chromeleon Studio - Chromatogram and Mass Spectra

## Slide 178: Chromatogram window
In the Data Processing pane in Chromeleon Studio, go to Data Processing Home tab, and ensure the icon is the only one selected.

## Slide 179: Chromatogram window
Highlighting the Chromatogram window brings up three tabs (Data Processing Home, Processing and Layout) as indicated by the red box.

## Slide 180: Removing Delimiters and Baseline from the Chromatogram
Go to the Layout tab.
Uncheck and in the Peak Characterization category.

## Slide 181: Removing Delimiters and Baseline from the Chromatogram
The Delimiters and Baseline are no longer present on the chromatogram.

## Slide 182: Chromatogram window: Seeing multiple chromatograms
You can see consecutive runs by pressing Shift (on the keyboard) and selecting the last run in the sequence of chromatograms you want to see.

## Slide 183: Chromatogram window: Seeing multiple chromatograms
In the Comparison section, switch from “Stacked” to “Overlaid”.
The chromatogram window should appear as shown above.

## Slide 184: Chromatogram window: Seeing multiple chromatograms
Next, deselect “Signal Offset” and “Time Offset”.
The chromatogram window should appear as shown above.

## Slide 185: Chromatogram window: Seeing multiple chromatograms
You can see specific runs in the Sequence by pressing Ctrl (on the keyboard) and selecting each run in the sequence of chromatograms you want to see.

## Slide 186: Chromatogram – Zoom in
To zoom in, left-click and highlight the region where you want to zoom in.

## Slide 187: Chromatogram – Zoom in
You should now see the zoomed in region of the chromatogram.
  The entire chromatogram is shown in an overview plot (red box).
NOTE: Left and right arrows pan across the chromatogram.

## Slide 188: Chromatogram – Overview plot
If you can’t see the overview plot, right-click and select Properties.

## Slide 189: Chromatogram – Overview plot
In the Properties Chromatogram window that appears, go to the “Interactive Settings” tab.
Click “Show overview plot” (red box).
  The Overview plot should now appear in the chromatogram window.

## Slide 190: Chromatogram – Zoom in Tips and Tricks
NOTE: You can also click in the middle of either the x-axis or y-axis to activate the Auto-zoom (red box).

## Slide 191: Chromatogram - Viewing data points Tips and Tricks
To see data points in the Chromatogram:
  Highlight the chromatogram.
  In the Layout tab, click on the “Plot Details” icon and select “Data Points”

## Slide 192: Chromatogram – Zoom out Tips and Tricks
NOTE: Left-clicking on the middle of the x-axis (red box) zooms out to the Full scale.

## Slide 193: Generating XICs
To generate XICs you can use MS AutoFilters (red box).

## Slide 194: Chromeleon Studio: Switching the View
NOTE: You can switch the view so that the Chromatogram and Mass Spectrum are side-by-side, by clicking Ctrl + Shift + F12 on your keyboard.

## Slide 195: Module 5: Building Processing Method
Module 5: Building Processing Method

## Slide 196: Adding the Processing Method window to the View Settings
To see the Processing Method window, click on the “ Processing Method” icon (red box) in the ribbon found under the Data Processing Home tab.

## Slide 197: The Processing Method window
The Processing Method window should now appear at the bottom.
It comprises of a number of tabs (red box).
In the next few Modules we will go through the relevant ones for MS/MS analysis.

## Slide 198: The Processing Method window Tips and Tricks
NOTE: You can add more or remove pages by right-clicking on one of the tabs and selecting “Page Selector”.

## Slide 199: The Processing Method window Tips and Tricks
In the Page Selector window, uncheck the Page you want to remove.
  In this example, we remove Peptide Table, MS Library Screening and Composite Scoring
Click “OK”.

## Slide 200: Saving the Processing Method View Setting
To save this view as a new View Setting, right-click on the currently selected View Setting and choose the “Save As” option.

## Slide 201: Saving the Processing Method View Setting
In the pop-up window that appears, input an Object name and click Save.

## Slide 202: Saving the Processing Method View Setting
There should now be two View Settings (Default and Processing Method) in the View Settings section.

## Slide 203: Setting Default View Setting for Chromeleon Studio
The Default View for your Chromeleon Studio is a property of the Sequence. To set the Default View, right-click on the Sequence in the Navigation pane, and choose Properties.

## Slide 204: Setting Default View Setting for Chromeleon Studio
The Sequence Properties window appears.
Click on the drop-down menu for “View Settings” and choose the Chromeleon Studio View template you want to use as default.

## Slide 205: Module 5.1: Data Processing - Processing Method / MS Component Table
Module 5.1: Data Processing - Processing Method / MS Component Table

## Slide 206: Importing new components
Click on the “MS Component Table” tab in the Processing Method window.
Next, click on the Processing Method tab in the ribbon and click icon.
  This is greyed out if the MS Component Table tab is selected.

## Slide 207: Importing new components
In the pop-up window that appears:
  Click the drop-down arrow (next to Browse) and select “MS Raw Data”.

## Slide 208: Importing new components
Browse to the location of the MS Raw Data you previously acquired.
Click Open.

## Slide 209: Importing new components
The Components should now appear in the “Compound Data Import” window.
Click Import.
NOTE: You have the option of manually selecting a few components from the list by checking only those components.

## Slide 210: Importing new components
When the import process is complete, the imported components should now appear in the MS Component Table.

## Slide 211: MS Components Table Tips & Tricks: Freeze Columns
Right-click on the top of the MS Components Table where the Column Names are located and choose the “Table Columns…” option.

## Slide 212: MS Components Table Tips & Tricks: Freeze Columns
In the Table Columns pop-up window that appears, input “3” for Frozen Columns (red box).
Click “OK”.

## Slide 213: MS Components Table Tips & Tricks: Freeze Columns
You should now be able to always see the first three columns (red box), even when you scroll to the MS Confirming Peak columns on the right.

## Slide 214: Module 5.2: Data Processing - MS Component Window
Module 5.2: Data Processing - MS Component Window

## Slide 215: MS Components window
In the Data Processing Home tab, deselect the Auto Filters icon, and select the MS Components icon.
  The MS Components window (red box) should now appear next to the Chromatogram.
  Make sure that the MS Quantitation channel is selected by default (blue box)

## Slide 216: MS Components window
Switch from a component to another by clicking on a different component in the Components list (red box).

## Slide 217: MS Components window
In the MS Settings tab, make sure that the Mass tolerance is fixed at 500 MMU and checked the Inhibit integration for TIC channel (red box).

## Slide 218: MS Components window
In the MS Detection tab, specify the Default Detection Algorithm, ICIS or Genesis for MS data (red box).

## Slide 219: MS Components window
In the MS Detection tab, specify the Default Detection Algorithm, ICIS or Genesis for MS data.

## Slide 220: Module 5.3: Data Processing – Calibration Levels
Module 5.3: Data Processing – Calibration Levels

## Slide 221: MS Components Table
You can create new concentration levels by right-clicking on the top of the MS Components Table where the Column Names are located and selecting Add Level.

## Slide 222: MS Components Table – Add Level
The Concentration Level box opens, choose the level name and click OK.

## Slide 223: Injection List – Level
After creating the different levels in the MS Component Table, specify the level for each injection the Injection List.
Or you can directly create new levels in the Injection List by clicking on Create New Level… (red box).

## Slide 224: MS Component Table
Specify the concentration unit in the Conc. Unit column (red box).

## Slide 225: MS Component Table
Specify the concentration unit in the Conc. Unit column (red box).
Press F9 to fill down.

## Slide 226: Module 5.4: Data Processing - Properties Component
Module 5.4: Data Processing - Properties Component

## Slide 227: MS Components Table – Properties Component
To open the Properties Component tab, double click in the MS Component Table.

## Slide 228: MS Components Table – Properties Component
Specify the retention time if it has not been well adjusted and the detection window.

## Slide 229: MS Components Table – Properties Component
In the Evaluation tab, specify whether if the compound is an Internal Standard or not (red box).

## Slide 230: MS Components Table – Properties Component
In the Evaluation tab, specify the calibration method: External (without Internal Standard) or Internal (select the Internal Standard associated with the compound).

## Slide 231: MS Components Table – Properties Component
In the Calibration tab, Calibration Type area: define the different calibration parameters (red box).

## Slide 232: MS Components Table – Properties Component
In the Extracted Ion Chromatograms tab, make sure that the MS Quantitation Peak and the MS Confirming Peak are well set.

## Slide 233: MS Components Table – Properties Component
If not, you can right-click on the MS Confirming Peak > Convert to MS Quantitation Peak.

## Slide 234: MS Components Table – Properties Component
In the Extracted Ion Chromatograms tab, the default integration settings can be changed by unchecking the Use Default MS Detection settings (red box).
This option let the possibility to use different algorithms for different components.

## Slide 235: MS Components Table – Properties Component
In the Chemical Details tab, add information about your components.

## Slide 236: MS Components Table – Properties Component
Switch from a component to an other with these options (red boxes).
Close the Properties Component table (blue box).

## Slide 237: Processing Method – Generate Ion Ratio Target
To define a target ion ratio as a confirmation criteria, go to Processing Method ribbon > XICs > Update (red box).

## Slide 238: Processing Method – Generate Target Ion Ratios
Check the components to use ion ratio confirmation (red box).
Choose the calculation method (blue box).
Click Apply and close the menu.

## Slide 239: Ion Ratio Confirmation
Now the ion ratio is enabled and the Target Ratio (%) has been calculated (red box).

## Slide 240: Ion Ratio Confirmation
Now the ion ratio is a confirmation criteria (red box).

## Slide 241: Module 5.5: Data Processing - MS Components
Module 5.5: Data Processing - MS Components

## Slide 242: MS Components
Click on the Chromatogram option to hide it and Interactive Results and Calibration Plot options to show them (red box).

## Slide 243: MS Components
Save the view setting as « Interactive Results ».

## Slide 244: MS Components
To add the Signal-to-Noise ratio on the MS Components view, right click on the chromatogram and click on Properties… (red box).

## Slide 245: MS Components
A table appears, go to the Peak Label tab.

## Slide 246: MS Components
In the Label identified peaks area, add « ; S/N= » (red box).
Click on the option (blue box).

## Slide 247: MS Components
In the Categories table, select Peak Results (red box).
In the Variables table, select Signal-to-Noise Ratio (blue box).

## Slide 248: MS Components
Click on the Parameters… buton to change the S/N calculation method (red box).

## Slide 249: MS Components
The Signal-to-Noise formula should appear in the Label identified peak field and the S/N value on the chromatogram (red boxes).

## Slide 250: Module 5.6: Data Processing – Calibration Plot
Module 5.6: Data Processing – Calibration Plot

## Slide 251: Calibration Plot
To remove the check standard points and the confidence interval from the calibration plot click on Check Standard Points and Confidence Interval (red boxes) in Calibration Plot Tools > Layout.

## Slide 252: Calibration Plot
To add the equation of the calibration curve and the R square: right click on the calibration plot > Properties… (red box).

## Slide 253: Calibration Plot
Go to Title and uncheck Use default (red box).
NOTE: the equation will be situated in the center, use left or right if preffered.

## Slide 254: Calibration Plot
Erase what was in the box and write « Y= ».
Click on (red box).

## Slide 255: Calibration Plot
Categories > Peak Calibration.
Variables > Calibration Coefficient.
Click OK.

## Slide 256: Calibration Plot
The calibration coefficient formula is now displayed.
NOTE: Y = a X² + b X+ c with {peak.calCoefficient(2)} = a / {peak.calCoefficient(1)} = b / peak.calCoefficient(0)} = c

## Slide 257: Calibration Plot
To add the R square click on (red box).

## Slide 258: Calibration Plot
Categories > Peak Calibration.
Variables > Coeff. Of Determination.
Click OK.

## Slide 259: Calibration Plot
Click Close.

## Slide 260: Calibration Plot
Now the equation of the calibration curve and the R square are displayed on the calibration plot (red box).
Don’t forget to save the view setting « Interactive Results » to conserve the changes (blue box).

## Slide 261: Calibration Plot
Zoom on the calibration plot to see the low concentration points.
The overview plot (red box) will appear, double click on it to unzoom.

## Slide 262: Calibration Plot
To disable a point from the calibration curve, select the injection (here CAL1-1), go to Calibration Plot Tools > Processing > Click on Disable Current (red box).
NOTE: This action will disable CAL1-1 only for the active compound.

## Slide 263: Module 5.7: Data Processing – Interactive Results
Module 5.7: Data Processing – Interactive Results

## Slide 264: Interactive Results – Summary (Unknows)
Summary (Unknows) table displays only the Unknow injections.

## Slide 265: Interactive Results – Summary (Check Standards)
Summary (Check Standards) table displays only the Check Standard injections.

## Slide 266: Interactive Results – Summary (Calibration)
Summary (Calibration) table displays only the Calibration Standard injections.

## Slide 267: Interactive Results – Peak Results
Peak Results table displays a summary of all the detected peaks.

## Slide 268: Interactive Results – Calibration
Calibration table displays a summary of the calibration for each compound detected in the selected injection.

## Slide 269: Module 5.8: Data Processing – Interactive Charts
Module 5.8: Data Processing – Interactive Charts

## Slide 270: Interactive Charts
Click on the Calibration option to hide it and Interactive Charts option to show it (red box).
Now the Interactive Charts is displayed.

## Slide 271: Interactive Charts
Save the view setting as « Charts ».

## Slide 272: Interactive Charts
To display the 3s lines, go to Interactive Chart Tools > Charts > 3s (red box)
NOTE: the chart is interactive, click on a data point to display the chromatogram.

## Slide 273: Interactive Charts
To add a Retention Time tab, go to Interactive Chart Tools > Charts > Retention Time (red box).

## Slide 274: Interactive Charts
To add a Retention Time tab, go to Interactive Chart Tools > Injection Chart > Click on the arrow (red box).
The Properties Interactive Charts appears.

## Slide 275: Interactive Charts
Unchecked the Blank injection type (red box).
Add a condition to filter the chart on the injection name (to remove the CAL0 point) (blue box).

## Slide 276: Interactive Charts
The biggest outliers have been removed.
Don’t forget to save the view setting « Charts » to conserve the changes (red box).

## Slide 277: Module 5.9: Data Processing - Using SmartLink in MS Components window
Module 5.9: Data Processing - Using SmartLink in MS Components window

## Slide 278: MS Components and Calibration Plot: SmartLink
You can see the XICs and the calibration plots for all the components simultaneously by clicking the SmartLink icon in the Data Processing Home ribbon (red box).

## Slide 279: Module 6: Report Templates
Module 6: Report Templates

## Slide 280: Creating a Report Template
To create a report, go to Report Designer.
Select the “Default MS Report” option (red box).
Click OK.

## Slide 281: Creating a Report Template
The new Template opens.

## Slide 282: Saving the new Report Template
You can save the new Template by going to the Report Templates section in the Navigation Pane (red box), right-clicking and selecting Save Changes.

## Slide 283: Saving the new Report Template
In the Save Modification pop-up window that appears, click Save.

## Slide 284: Renaming the new Report Template
You can also rename the new Template by going to the Report Templates section in the Navigation Pane (red box), right-clicking and selecting Rename.

## Slide 285: Renaming the new Report Template
In the pop-up window that appears, input a Name and click “OK”.

## Slide 286: Renaming the new Report Template
The Template with its new name should now appear in the Report Templates section of the Navigation Pane (red box).
NOTE: Clicking the icon that appears saves the changes made to the Template.

## Slide 287: Deleting a sheet
To delete a sheet from the template, go to the sheet, in the ribbon select the arrow next to Delete and select Delete Sheet (red box).

## Slide 288: Deleting a sheet
Select Yes.

## Slide 289: Sequence Overview Sheet
The sequence overview gives information from the injection list.

## Slide 290: Integration Sheet
The integration table summaries information regarding the compounds found in the sample.
We can filter the table to conserve only the target compounds (not the internal standards).

## Slide 291: Integration Sheet
Go to the Integration Results table > Table Tools > Integration Table > Click on the arrow (red box).
A pop-up window opens.

## Slide 292: Integration Sheet
Check “Only include peaks that match the following”
Click on > Component > Standard Method > Click OK.

## Slide 293: Integration Sheet
Select “Does not contain” and “ISTD Internal”.
Click Close

## Slide 294: Integration Sheet
Now only the target compounds are displayed.

## Slide 295: Create a Check Standard sheet
To change the sheet name double-click on the Component Sequence Summary (blue box).
A pop-up window opens (red box)
Change the name and click OK.

## Slide 296: Create a Check Standard sheet
To change top page name, go to cell A1 and change the name in the writing area (red box).
Click Enter.

## Slide 297: Create a Check Standard sheet
The name has changed.
To filter out injections types go to Table Tools > Summary Table > in the Include Injection Types uncheck everything except Check Standard (red box).

## Slide 298: Create a Check Standard sheet
To add a more specific filtering condition you can click on the arrow (red box).
Check “Only include injections that match the following conditions” (blue box).
In this example I will keep only injections that contain “CTL” in their name.

## Slide 299: Create a Check Standard sheet

## Slide 300: Create a Check Standard sheet
Now only the CTL Check Standards are displayed.

## Slide 301: Create a Check Standard sheet
To change the formula of the last column, go to the column header.
Table Tools > Summary Table > Column Properties > Click .
A pop-up window opens and select Catergories > Peak Results and Variables > Amount Deviation
Click OK.

## Slide 302: Create a Check Standard sheet
In the formula peak.amount_deviation("abs") change « abs » by « rel »
Click OK.

## Slide 303: Create a Check Standard sheet
Now the relative difference between the theoretical concentration and the calculated concentration is displayed.

## Slide 304: Create a Check Standard sheet
You can save the report by clicking Save (red box).
NOTE: you can create as many reports as you want within a sequence.

## Slide 305: Module 7: Creating an eWorkflow
Module 7: Creating an eWorkflow

## Slide 306: Creating an eWorkflow
In the eWorkflow category bar, go to the Home ribbon, click “Create” and select “eWorkflow…”.

## Slide 307: eWorkflow: The eWorkflow General tab
In the eWorkflow Editor pop-up window that appears, go to the eWorkflow General tab, select HPLC under the eWorkflow Type.

## Slide 308: Add Instruments
In the Instruments section, click “Add” (red box).

## Slide 309: Add Instruments
In the Select Instrument window, choose the instrument and click “OK”.
NOTE: You can select multiple instruments.

## Slide 310: Add Methods
In the Instruments section, click “Add” (red box).

## Slide 311: Add Methods
In the Add Chromeleon Method window, choose a Sequence and click “Add” (red box).

## Slide 312: Add Methods
Next, select Instrument Methods, Processing Method, Report Template & View Setting, then click “Add” (red box).
NOTE: You can browse in multiple items simultaneously.

## Slide 313: Add Methods
In the Save As window that appears there’s a choice between two options:
  “Save complete copy” ensures that each time the eWorkflow is executed the same exact copy of the item is used in the resulting Sequence.
  “Save only links” ensures that each time the eWorkflow is executed the most current version of the item is used in the resulting Sequence.

## Slide 314: Add Attachments
In the Attachments section, click “Add” (red box).

## Slide 315: Add Attachments
Attach laboratory SOP’s or other items such as Reports to the eWorkflow.
NOTE: Doing this copies the item into the resulting Sequence providing the user with quick access to the document.

## Slide 316: Add Attachments
Again, in the Save As window that appears there’s a choice between two options:
  “Save complete copy” ensures that each time the eWorkflow is executed the same exact copy of the item is used in the resulting Sequence.
  “Save only links” ensures that each time the eWorkflow is executed the most current version of the item is used in the resulting Sequence.

## Slide 317: eWorkflow: The eWorkflow General tab
The eWorkflow General tab should now appear as shown above.
NOTE: the eWorkflow State (redbox) is useful when using different user accounts > a User Role can be defined to launch only Approved eWorkflows.

## Slide 318: eWorkflow: The Sequence General tab
Left-click on the Sequence General tab (red arrow).
This tab should look like that shown above.

## Slide 319: eWorkflow: The Sequence General tab
In this tab you can set default Report Templates, View Settings and Channels (red box).

## Slide 320: Choose the Data Vault
You can browse in the Data Vault (red arrow).
NOTE: The Current Data Vault is the Chromeleon Local (blue arrow). By default, there is only one Data Vault.

## Slide 321: Input a Path
You can then change the path by adding a formula in the Path bar (blue box) so that it creates folders and sub-folders with names extracted from details related to the run.
NOTE: This makes for easy archiving, especially when running multiple instruments over a network.

## Slide 322: Input a Path
To do this, first input “AminoAcids/” in the Path bar.

## Slide 323: Input a Path
Next click the icon next to the Path bar to bring up the Insert Formula window.

## Slide 324: Input a Path
Select “eWorkflow” in the Categories section and “Instrument Used” in the Variables section.

## Slide 325: Input a Path
Click “OK” and input “/” at the end of the formula.

## Slide 326: Input a Path
Next, click the icon next to the Path bar to bring up the Insert Formula window.
Select “eWorkflow” in the Categories section and “Description” in the Variables section.

## Slide 327: Input a Path
Click “OK” and input “/” at the end of the formula.

## Slide 328: Input a Path
Next, click the icon next to the Path bar to bring up the Insert Formula window.
Select “General” in the Categories section and “Current Time” in the Variables section.
In the Format bar input “YYYY”.

## Slide 329: Input a Path
Click “OK” and input “/” at the end of the formula.

## Slide 330: Input a Path
Next, click the icon next to the Path bar to bring up the Insert Formula window.
Select “General” in the Categories section and “Current Time” in the Variables section.
In the Format bar input “mm mmm”.

## Slide 331: Input a Path
Click “OK”.
The Vault and Path should now appear as shown above (red box).

## Slide 332: Input a Sequence Name
To add a Sequence Name, click the icon next to the Sequence Name bar.
In the Insert Formula window that appears, select “General” in the Categories section and “Current Time” in the Variables section.

## Slide 333: eWorkflow: The Sequence General tab
Click “OK”.
The Sequence General tab should now appear as shown above.

## Slide 334: eWorkflow: The Sequence General tab
In the Notifications section, select “Enable notifications”.
Also deselect “Submit electronic signatures” and “Smart Startup” if you don’t require them.

## Slide 335: eWorkflow: The Sequence Features tab
In the Sequence Features tab you have the option of seeings changes made in the Chromeleon Studio immediately (“Automatically”) or when you update it (“On Demand”).

## Slide 336: eWorkflow: The Sequence Features tab
NOTE: Whenever the “On Demand” option is selected, you will see the yellow bar below the ribbon in the Chromeleon Studio whenever you make a change.

## Slide 337: eWorkflow: The Sequence Layout tab
The layout of the Injection List is setup in the Sequence Layout tab (red arrow).
NOTE: You can designate the frequency of periodic samples in the Max Samples per Bracket field and limit how many samples are run in a Sequence in the Max Brackets per Sequence (blue arrows).
Sequence Header Block: Enter Injections that always appear at the beginning of a Sequence, such as Blank Injections.
Sample Block: Enter a single line to represent the sample injections.
Bracket Block: Enter samples that should be inserted at the frequency set in the Max Samples per Bracket field.
Sequence Footer Block: Enter those Injections that are run at the end of every Sequence, such as a final Blank.

## Slide 338: eWorkflow: The Sequence Layout tab
Change the Max Brackets per Sequence from the default of “1” to “99” (blue arrow).
Next, determine the Injection List Layout, and finally build the Injection List on the Sequence Layout tab by clicking the “New eWorkflow Injection Template” hyperlink in the appropriate block.

## Slide 339: Adding a Blank Injection to the Sequence Header Block
In this example, we will be adding two Blank injections to the Sequence Header.
Clicking the “New eWorkflow Injection Template” hyperlink brings up a window.
Enter a name, click the > icon and select the “#r” option so that each injection is numbered.

## Slide 340: Adding a Blank Injection to the Sequence Header Block
Click “Add”.
A Blank injection should now appear in the Sequence Header block with an Instrument and Processing Method.

## Slide 341: Adding a Blank Injection to the Sequence Header Block
Change the No. of inj. To “2”.
NOTE: You can preview the Injection List in the Sequence Preview Section at the bottom of the editor window.

## Slide 342: Adding Calibration Standards to the Sequence Header Block
Next, click on the “New eWorkflow Injection Template” hyperlink in the Sequence Header to add the calibration standards.
Input “CAL#r-1” for the Injection Template Name and
Click “Add”.

## Slide 343: Adding Calibration Standards to the Sequence Header Block
In No. of Inj. Column add as many injections as needed.
Change Type from Blank to Calibration Standard.

## Slide 344: Adding samples in the Sample Block
Click on the “New eWorkflow Injection Template” hyperlink in the Sample Block to add injection for your Sample.
Input “Patient-#r” for the Injection Template Name and
Click “Add”.

## Slide 345: Adding samples in the Sample Block
Now the sample is added.

## Slide 346: Adding samples in the Bracket Block
Click on the “New eWorkflow Injection Template” hyperlink in the Bracket Block to add injection for your sample.
Input “CTL#r” for the Injection Template Name and
Click “Add”.

## Slide 347: Adding samples in the Bracket Block
Define the CTL#r sample as a Check Standard and No. of Inj. to 2.
Check Use Bracket Block after Sequence Header (red box).

## Slide 348: Adding samples in the Bracket Block
Define the CTL#r sample as a Check Standard and No. of Inj. to 2.
Check Use Bracket Block after Sequence Header (red box).

## Slide 349: Adding a Shutdown method in the Sequence Footer Block
Click on the “New eWorkflow Injection Template” hyperlink in the Sequence Footer to add a calibration points at the end of the sequence and a blank injection.
NOTE: Change the Instrument Method to the “Standby” method for the blank sample.

## Slide 350: Assigning an Instrument
You can assign an Instrument by clicking on the dropdown menu (red arrow) and assigning an instrument.

## Slide 351: Assigning an Instrument
In the Select Instrument pop-up window that appears, select the instrument you want to use, then click “OK”.

## Slide 352: Assigning Sample Positions
Once an Instrument is assigned, the positions will automatically be populated.
  The Position changes from “n.a.” to positions that match the autosampler being used (red arrow).
NOTE: You can change the Sample start position to match the location of your sample vials (blue arrow).

## Slide 353: Saving the eWorkflow
You can save the eWorkflow by clicking on File > Save As... (red box).

## Slide 354: Saving the eWorkflow
Input “eWorkflow_AminoAcids” for the Name in the Save eWorkflow window that appears.
Click “Save”.

## Slide 355: Launching the eWorkflow
Close the eWorkflow Editor.
The eWorkflow should now be listed in the Navigation Pane under the eWorkflow Category Bar in Chromeleon.
NOTE: You can run the eWorkflow by clicking on “Launch” (red box).

## Slide 356: Launching the eWorkflow
The eWorkflow wizard should now appear. In this example, there are 20 samples (red box).

## Slide 357: Viewing the Sequence generated by the eWorkflow
Click Finish.
In the Data Category bar, if you used the formulas described in the Slides above, a folder named “AminoAcids” with subfolders named after the instrument (Vanquish-Altis), year (2020), month (03 Mar) and Sequence Name (27_Mar_2020 09_36) should be created.

## Slide 358: Viewing the Sequence generated by the eWorkflow
The injections should now appear in the Injection list for the Sequence you just created.
NOTE: You can now change the Sample Name, Positions, Volume and any other parameters as needed.

## Slide 359: Viewing the Sequence generated by the eWorkflow
To change the sample name, highlight the samples, click F9.
A pop up window opens.

## Slide 360: Viewing the Sequence generated by the eWorkflow
Change Start from 0 to 1.
Use pattern : Patient-#u.
Click OK.

## Slide 361: Viewing the Sequence generated by the eWorkflow
Now the samples are named.

## Slide 362: Starting the Sequence run
Click Start to run the sequence.

## Slide 363: Module 8: Exporting and Importing Data
Module 8: Exporting and Importing Data

## Slide 364: Exporting Data
The File > Send To… option creates cmdx data files for export of data.
  Automatically checks “Include 3D (UV) and MS raw data”.*
    If this is unchecked, the exported data only includes methods and reports (no MS data).
  “Include Versions” means all saved versions will be included in the cmdx file.
* Only applicable for Chromeleon 7.2 SR3 and later.

## Slide 365: Exporting Data
To open the files in a different Thermo software (Freestyle, TraceFinder) you must export the data in .raw file format.
  To do this, right-click on the Sequence and choose “Export” option.

## Slide 366: Exporting Data
Choose the .raw option.
  Click on the icon to dictate how the file is named (See inbox).

## Slide 367: Exporting Data
Choose destination path by browsing folder location in the Parent Folder section.
You can also use a formula to create a Sub-Folder.
Click OK.

## Slide 368: Exporting Data
The .raw file should now appear in the folder you just created.
NOTE: The sub-folder that is created is named “20_Mar_2020 15_03” (red arrow).

## Slide 369: Importing Data (Chromeleon Data)
Select the folder in the Navigation pane where you want to add the data.
Next, go to File > Import, and select “Chromeleon Data…”.
In the Open window that appears, select the .cmbx file you need to import and click “Open”.

## Slide 370: Importing Data (Chromeleon Data)
In the Import window that appears, you can import the .cmbx file either to the Original location or another location.
Click “Start”.
The sequence should start importing, as indicated by the progress bar.

## Slide 371: Importing Data (Bulk Import)
You can also import data that was acquired outside of Chromeleon.
Select the folder in the Navigation pane where you want to add the data.
Next, go to File > Import, and select “Bulk Import…”.

## Slide 372: Importing Data (Bulk Import)
You can either drag and drop the files and folders into the left pane or click one of the icons at the top.
Clicking the icon, brings up a pop-up window, where you can select the files you want to import into Chromeleon.
Next, click “OK”.
Notes: one of the icons at the top.

## Slide 373: Importing Data (Bulk Import)
The files should now appear in the left pane.
You can add more files by clicking the icon and Browsing in more data.
When done, click “Import”.
The sequence should start importing.
  Progress bar sits at either 0% or 100%.
Notes: one of the icons at the top.

## Slide 374: Importing Data (Bulk Import)
When the import is complete; once successful, you should see a message similar to the one shown above.
Click “Close”.
Go to the Data Category bar, you should now see the Sequence (red arrow) with the data you just imported.
Click “Convert” on the blue task bar at the top.

## Slide 375: Importing Data (Bulk Import)
Click the Save icon (red box) to save the imported Sequence.

## Slide 376: Appendix

## Slide 377: Useful Function keys (F1, F2, F4/Shift+F4 & F5)
https://www.youtube.com/watch?v=ejnopr7NLF4&list=UUiJkt03U-392FA-CdgqoSWw

## Slide 378: MS AutoFilters
https://www.youtube.com/watch?v=1cFMBkjQPGo&index=2&list=UUiJkt03U-392FA-CdgqoSWw

## Slide 379: Channels
https://www.youtube.com/watch?v=ZrteLQoz3hM&index=7&list=UUiJkt03U-392FA-CdgqoSWw

## Slide 380: Custom Reports in Chromeleon Part 1
https://www.youtube.com/watch?v=m6yUF2CUMpo&list=UUiJkt03U-392FA-CdgqoSWw&index=10

## Slide 381: Custom Reports in Chromeleon Part 2
https://www.youtube.com/watch?v=uQQcfeMseTk&list=UUiJkt03U-392FA-CdgqoSWw&index=9
FAIL

## Slide 382: Custom Reports in Chromeleon Part 3
https://www.youtube.com/watch?v=NdTQevHAT4c&index=8&list=UUiJkt03U-392FA-CdgqoSWw

## Slide 383: How to remove Charlie watermark from Chromeleon reports
https://www.youtube.com/watch?v=cVc_P1lHX5s&list=UUiJkt03U-392FA-CdgqoSWw&index=14